Piezotactile Stimulator PTS-C2 - Instructions for Use

General description

The PTS stimulator system consists of two piezoelectric stimulator heads, a 2-channel amplifier with built-in oscillators, a battery pack and battery charger. Optionally there may be an RFI filter.

IMPORTANT - Precautions

The amplifier generates high voltages to drive the piezoelectric wafers. Do not attempt to repair or disassemble any part of the system. If a probe or cable becomes damaged stop using it immediately.

The piezoelectric wafer is brittle and will be destroyed by a force of more than about 500g applied to the probe. This is moderate finger pressure. Subjects and experimenters should be asked not to press on the probe.

The wafer has a mechanical resonance at about 180Hz. Excessive excursion at this frequency may damage the wafer. For this reason do not drive the device at 160 - 200Hz unless it is applied to a mechanical load which will damp the resonance, such as the skin. When working in this frequency range if possible test the amplitude starting with a low amplitude stimulus and working up. Avoid driving the probe at full amplitude.

Note that when using an external waveform source some waveforms may have strong harmonics in the 180Hz range which may overdrive the wafer. For instance a 60Hz square wave has a strong 3rd harmonic at 180Hz. For this reason using waveforms with rapid transients such as square waves, steps, pulses or logic signals should be avoided.

Driving the piezo wafer at high amplitude with very low frequency (< 1Hz or DC) signals may cause the piezoelectric material to depole and decrease the efficiency of the device, so this should be avoided if possible. Note that logic signals always have a DC component, which is another reason not to use logic signals as a signal source.

The stimulator head

The stimulator head consists of a piezoelectric bimorph element (wafer) housed in a ceramic case. Attached to the wafer is an adjustable 8mm probe which protrudes through a 10mm hole in the case, which acts as a fixed surround. Indentation of the probe may be adjusted using a small tool provided. One turn of the probe changes the indentation by 1mm. The probe may be removed for cleaning.

Also housed in the stimulator head is an optical displacement transducer which measures the position of the wafer.

The ceramic case has a soft PVC cover fitted. This may be removed if desired but note that the case is quite brittle and may break if dropped on a hard surface.

The amplifier unit

The amplifier contains two completely independent identical channels. If a channel is not in use it should be switched off to conserve battery life. Each channel contains a high voltage drive amplifier to drive the piezoelectric element, a wide-range sine-wave oscillator to provide a signal source, and the detector electronics for the displacement transducer.

The piezo amplifier

The amplifier converts low-voltage signals into the high-voltage signals to drive the piezo element. Its only control is the amplitude control on the front panel.

The oscillator

The oscillator generates a stable sine-wave of between 0.45 and 600 Hz, in 3 switched ranges. The oscillator is connected to the amplifier when the Source switch is set to "On" or "Gated".

In the On mode the oscillator runs continuously. In the Gated mode the oscillator is controlled by a logic-level (0 - 5V) signal connected to the Gate input on the rear panel. A logic 0 on the Gate input stops the oscillator, a logic 1 starts the oscillator. The oscillator always starts and ends at the zero-crossing point of the waveform. This means that if the oscillator is gated on by a positive-going pulse there will always be a whole number of cycles in the resulting sine-wave burst. This in turn means that the sine wave burst may continue after the gate signal has gone to logic 0 for a period equal to 1/f, where f is the frequency of the oscillator.

A square wave of frequency equal to the oscillator frequency is available at the Freq output on the rear panel. This is intended for connecting to a frequency meter in order to set the frequency to some required value.

External signals

The amplifier may be fed with a signal from an external device such as a waveform generator or PC card with analog output. This signal should be connected to the Ext In connector on the rear panel. The Source switch should be set to "Ext". In this mode of operation the amplitude control will still work but the gate input will not.

Recommended amplitude of the external signal for full output swing of the amplifier is +/- 5V.

The displacement transducer

The displacement transducer produces a voltage output at the DT Out connector on the rear panel. The voltage varies linearly with the instantaneous displacement of the probe, with a ratio 10mV = 1 micron or 1V = 0.1mm. The transducer is intended for use when measuring vibrotactile thresholds using the PTS, or when it is desired to deliver a stimulus of known amplitude.

Note that the transducer electronics in the amplifier is calibrated for the transducer in the corresponding stimulus head, which means that if a different probe head is used the output of the transducer will not be accurate.

The battery

The battery is a 2 x 12V lead-acid type. Battery life will depend on number of channels used, stimulus duration, intensity and frequency, but a fully charged battery should provide enough power for a full day of typical use. The battery is charged by connecting its output terminal to the 24V battery charger provided. The charger will deliver a heavy charge current until the battery reaches full voltage, followed by a trickle charge to keep the battery at full capacity. There is no maximum charge time; it may be left on charge indefinitely.

Completely discharging the battery frequently may reduce its capacity, so this should be avoided. Do not leave the battery fully discharged for any length of time as this may permanently damage it. Always store the battery fully charged, and recharge it periodically even if it is not used.