Dancer Design

INNOVATION IN ELECTRONICS

Custom solutions for Neuroscience and Psychology Research

Unit 12, Catapult Too, Parade Street, St. Helens, Merseyside, WA10 1LX

Tel: 01744 610060  Mob: 07743 089056  Email: chris@dancerdesign.co.uk  Web: www.dancerdesign.co.uk

"Tactor" Miniature Stimulators

Dancer Design tactors are miniature vibrotactile electromagnetic solenoid-type stimulators. They are designed to be attached to the skin using adhesive rings, but may also be incorporated into objects or clothing.

The tactor can deliver a strong tactile stimulus at frequencies ranging from zero (static indentation) to 300 Hz. The vibration amplitude varies linearly with voltage to a good approximation. These characteristics allow for the delivery of qualitatively and quantitatively different stimuli, which may be useful in conveying information to a person.

Features

- Small size: 18mm diameter, 12mm high

- Light weight: 5.4 grams

- Design optimised for efficiency using finite element analysis software

- Flat probe and fixed surround optimally stimulates RA 1 receptors for localised stimulation

- Can be driven by standard audio amplifiers or Dancer Design Tactamp

- Wide frequency range: 0 – 300Hz

- Surfaces washable in alcohol for disinfection

- Screened design means low EM emission. Compatible with EEG.

Electrical and mechanical properties

- DC resistance: 18Ω

- Recommended maximum drive voltage: 6V peak sine wave, 6V DC

- Maximum free extension of probe at 6V DC: ~2mm

- Maximum static force delivered by probe at 6V DC: 0.35N

Materials

- Base: FR4 glass reinforced epoxy laminate

- Probe: Nylon

- Cover: Anodised aluminium

- Connector pins: Tin plated brass

- Adhesive attachment ring: disposable hypoallergenic adhesive tape

Force / frequency plot

Probe attached to rigid strain gauge. Drive voltage: sine wave, 6V peak.

Force output / voltage input characteristics at 20Hz

For further information, advice on use of tactors and custom drive solutions, contact Dancer Design: chris@dancerdesign.co.uk